## Instructions
QCnomics — Measurement Uncertainty (MU) workbook
Method: NABL 112A, Cl. 6(i)(iii) — MU is estimated for each quantitative test from at least 6 months of internal QC (IQC) data, per QC level.
Formula: CV% = SD ÷ Mean × 100; MU% = 1.96 × CV% (expanded uncertainty, about 95% confidence); MU at the mean = ± 1.96 × SD.
How to use
1. Open a "Test 1" … "Test 15" sheet (one sheet per test). Fill the yellow cells: test name, unit, analyzer, lab name, QC lot.
2. Paste the IQC results of each level into its Value column (Level 1: columns A–B, Level 2: C–D, Level 3: E–F). Add the dates if you have them.
3. n, Mean, SD, CV%, MU% and the ± value at the mean are calculated at the top of the sheet. "Check" confirms 6 months of data.
4. The "MU Register" sheet collects every test and level in one table — print it or attach it to your MU SOP / register.
5. Use only accepted IQC results. Exclude values rejected under your Westgard rules and results from a known error that was corrected.
6. Keep one QC lot per level where possible. If the lot changed in the period, start a second sheet or note it in the lot cell.
7. The "Example" sheet shows the expected format with sample data. Need more test sheets? Right-click a "Test" sheet → Move or Copy → Create a copy.
Colour key: yellow = you type here; green = MU results; white = calculated (do not edit).
Reporting a patient result: result ± (MU% × result ÷ 100), using the level closest to the result.
Online calculator: https://qcnomics.com/qc-calculator.html#mu-calculator · QCnomics computes MU automatically from your IQC data.

## MU Register
MU Register — expanded uncertainty from IQC (NABL 112A: MU = 1.96 × CV%)
Filled in automatically from the "Test 1" … "Test 15" sheets. Do not type here.
Sheet | Test / analyte | Unit | Analyzer / method | QC level | QC material / lot | n | Mean | SD | CV % | MU % | MU at mean (±) | IQC from | IQC to | Months | Check

## Example
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii) [EXAMPLE — sample data]
Test / analyte | Glucose (example)
Unit | mg/dL
Analyzer / method | Hexokinase
Lab name | Your laboratory
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot | L1 lot 45671 | L2 lot 45672
No. of results (n) | 130 | 130 | 0
Mean | 94.9831 | 240.5023
SD | 1.9796 | 4.9623
CV % | 2.0842 | 2.0633
MU % (1.96 × CV%) | 4.085 | 4.0441
MU at the mean (± in unit) | 3.8801 | 9.7262
IQC period: first date | 2026-01-01 | 2026-01-01
IQC period: last date | 2026-07-01 | 2026-07-01
Period (months) | 5.9461 | 5.9461
Check | OK — meets NABL 112A | OK — meets NABL 112A
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value
2026-01-01 | 94.5 | 2026-01-01 | 242.6
2026-01-02 | 96.1 | 2026-01-02 | 240
2026-01-04 | 94.5 | 2026-01-04 | 243.5
2026-01-05 | 94.3 | 2026-01-05 | 242.6
2026-01-07 | 93 | 2026-01-07 | 249.2
2026-01-08 | 94.6 | 2026-01-08 | 241.5
2026-01-09 | 97.3 | 2026-01-09 | 238
2026-01-11 | 95.9 | 2026-01-11 | 238.3
2026-01-12 | 97.2 | 2026-01-12 | 239.9
2026-01-14 | 95.5 | 2026-01-14 | 244.2
2026-01-15 | 95.8 | 2026-01-15 | 238.5
2026-01-16 | 95.4 | 2026-01-16 | 241.8
2026-01-18 | 91.5 | 2026-01-18 | 248.5
2026-01-19 | 96.8 | 2026-01-19 | 228.2
2026-01-21 | 96.1 | 2026-01-21 | 234.8
2026-01-22 | 96 | 2026-01-22 | 241.1
2026-01-23 | 91.4 | 2026-01-23 | 241.8
2026-01-25 | 91.3 | 2026-01-25 | 241.1
2026-01-26 | 93.1 | 2026-01-26 | 238
2026-01-28 | 94 | 2026-01-28 | 243
2026-01-29 | 95.6 | 2026-01-29 | 241.3
2026-01-30 | 94.9 | 2026-01-30 | 237.6
2026-02-01 | 96.1 | 2026-02-01 | 251.2
2026-02-02 | 93.7 | 2026-02-02 | 241.6
2026-02-04 | 95.6 | 2026-02-04 | 237.5
2026-02-05 | 95.8 | 2026-02-05 | 239.5
2026-02-06 | 93.6 | 2026-02-06 | 239
2026-02-08 | 98.6 | 2026-02-08 | 239.7
2026-02-09 | 96.2 | 2026-02-09 | 227.5
2026-02-11 | 97.5 | 2026-02-11 | 237.8
2026-02-12 | 93.7 | 2026-02-12 | 244.6
2026-02-13 | 93.4 | 2026-02-13 | 234.6
2026-02-15 | 94.3 | 2026-02-15 | 239.7
2026-02-16 | 94.8 | 2026-02-16 | 244.4
2026-02-18 | 96.3 | 2026-02-18 | 243.9
2026-02-19 | 95.5 | 2026-02-19 | 246.9
2026-02-20 | 94.1 | 2026-02-20 | 232.2
2026-02-22 | 93 | 2026-02-22 | 238.4
2026-02-23 | 93.9 | 2026-02-23 | 238.4
2026-02-25 | 97.6 | 2026-02-25 | 242.9
2026-02-26 | 93.3 | 2026-02-26 | 245
2026-02-27 | 95.5 | 2026-02-27 | 227.7
2026-03-01 | 95.9 | 2026-03-01 | 245
2026-03-02 | 91.9 | 2026-03-02 | 233.3
2026-03-04 | 95.1 | 2026-03-04 | 243.1
2026-03-05 | 97.7 | 2026-03-05 | 233.1
2026-03-06 | 90.8 | 2026-03-06 | 240.8
2026-03-08 | 94.3 | 2026-03-08 | 245.5
2026-03-09 | 94.8 | 2026-03-09 | 239.3
2026-03-11 | 93.3 | 2026-03-11 | 240.9
2026-03-12 | 96 | 2026-03-12 | 243.7
2026-03-13 | 94.9 | 2026-03-13 | 240.7
2026-03-15 | 91.9 | 2026-03-15 | 239.6
2026-03-16 | 96.7 | 2026-03-16 | 247.1
2026-03-18 | 96.4 | 2026-03-18 | 244.8
2026-03-19 | 97 | 2026-03-19 | 238.6
2026-03-20 | 98 | 2026-03-20 | 252.6
2026-03-22 | 95.8 | 2026-03-22 | 234.7
2026-03-23 | 95.3 | 2026-03-23 | 244.2
2026-03-25 | 92.3 | 2026-03-25 | 238.8
2026-03-26 | 96.3 | 2026-03-26 | 240.6
2026-03-27 | 93.7 | 2026-03-27 | 243.2
2026-03-29 | 94 | 2026-03-29 | 241
2026-03-30 | 92.3 | 2026-03-30 | 242.9
2026-04-01 | 93 | 2026-04-01 | 233
2026-04-02 | 93.9 | 2026-04-02 | 233.1
2026-04-03 | 97.7 | 2026-04-03 | 242.8
2026-04-05 | 90.7 | 2026-04-05 | 235.6
2026-04-06 | 91.9 | 2026-04-06 | 235.3
2026-04-08 | 95.5 | 2026-04-08 | 233.2
2026-04-09 | 98 | 2026-04-09 | 245.8
2026-04-10 | 96.2 | 2026-04-10 | 243.4
2026-04-12 | 91 | 2026-04-12 | 246.8
2026-04-13 | 89.7 | 2026-04-13 | 235.7
2026-04-15 | 95.8 | 2026-04-15 | 240
2026-04-16 | 93.5 | 2026-04-16 | 234.8
2026-04-17 | 92.6 | 2026-04-17 | 243.5
2026-04-19 | 97.1 | 2026-04-19 | 247.3
2026-04-20 | 97.3 | 2026-04-20 | 235.9
2026-04-22 | 95.3 | 2026-04-22 | 247.2
2026-04-23 | 95.5 | 2026-04-23 | 244.5
2026-04-24 | 95.9 | 2026-04-24 | 239.2
2026-04-26 | 98.3 | 2026-04-26 | 230.9
2026-04-27 | 96.3 | 2026-04-27 | 246.5
2026-04-29 | 96.1 | 2026-04-29 | 239.6
2026-04-30 | 96.2 | 2026-04-30 | 237.2
2026-05-01 | 91.7 | 2026-05-01 | 241.8
2026-05-03 | 97.7 | 2026-05-03 | 241.9
2026-05-04 | 97 | 2026-05-04 | 246.9
2026-05-06 | 96.1 | 2026-05-06 | 235.3
2026-05-07 | 90.9 | 2026-05-07 | 245.2
2026-05-08 | 93.7 | 2026-05-08 | 246.8
2026-05-10 | 96.8 | 2026-05-10 | 246.7
2026-05-11 | 91.2 | 2026-05-11 | 239.2
2026-05-13 | 94.6 | 2026-05-13 | 236.6
2026-05-14 | 97.1 | 2026-05-14 | 244.7
2026-05-15 | 92.2 | 2026-05-15 | 240.5
2026-05-17 | 98.4 | 2026-05-17 | 240.6
2026-05-18 | 96.2 | 2026-05-18 | 246.6
2026-05-20 | 94.7 | 2026-05-20 | 238.8
2026-05-21 | 95.7 | 2026-05-21 | 229.4
2026-05-22 | 96.4 | 2026-05-22 | 238.2
2026-05-24 | 95.3 | 2026-05-24 | 231.5
2026-05-25 | 97.4 | 2026-05-25 | 243.8
2026-05-27 | 93.6 | 2026-05-27 | 241.5
2026-05-28 | 94.1 | 2026-05-28 | 237.2
2026-05-29 | 97.2 | 2026-05-29 | 240
2026-05-31 | 95.1 | 2026-05-31 | 243.8
2026-06-01 | 93.2 | 2026-06-01 | 240.4
2026-06-03 | 97 | 2026-06-03 | 246.1
2026-06-04 | 98.1 | 2026-06-04 | 239.7
2026-06-05 | 94.1 | 2026-06-05 | 244.8
2026-06-07 | 92.1 | 2026-06-07 | 246.9
2026-06-08 | 94.7 | 2026-06-08 | 247.4
2026-06-10 | 94.7 | 2026-06-10 | 236.9
2026-06-11 | 94.4 | 2026-06-11 | 244
2026-06-12 | 98 | 2026-06-12 | 231.4
2026-06-14 | 92.8 | 2026-06-14 | 235
2026-06-15 | 97.6 | 2026-06-15 | 231
2026-06-17 | 92.3 | 2026-06-17 | 244.9
2026-06-18 | 93.3 | 2026-06-18 | 234.3
2026-06-19 | 96.3 | 2026-06-19 | 239.9
2026-06-21 | 97.4 | 2026-06-21 | 239.1
2026-06-22 | 96.8 | 2026-06-22 | 239.9
2026-06-24 | 95.7 | 2026-06-24 | 237.3
2026-06-25 | 95.3 | 2026-06-25 | 241.1
2026-06-26 | 95.3 | 2026-06-26 | 248.2
2026-06-28 | 96.2 | 2026-06-28 | 240.2
2026-06-29 | 94.6 | 2026-06-29 | 242.4
2026-07-01 | 95.6 | 2026-07-01 | 244.6

## Test 1
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 2
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 3
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 4
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 5
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 6
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 7
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 8
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 9
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 10
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 11
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 12
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 13
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 14
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value

## Test 15
Measurement Uncertainty from IQC — NABL 112A Cl. 6(i)(iii)
Test / analyte
Unit
Analyzer / method
Lab name
Results (calculated automatically — do not type here)
Parameter | Level 1 | Level 2 | Level 3
QC material / lot
No. of results (n) | 0 | 0 | 0
Mean
SD
CV %
MU % (1.96 × CV%)
MU at the mean (± in unit)
IQC period: first date
IQC period: last date
Period (months)
Check
Type or paste IQC results into the yellow cells below — one result per row, each level in its own pair of columns. Dates are optional but needed for the 6-month check. Up to 1000 results per level.
Level 1 |  | Level 2 |  | Level 3
Date | Value | Date | Value | Date | Value